## Sheet1
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol | Acceptance Criteria | Comments
5.11.1 | Patient Expectations | 5.1.2 | Can be worn up to 12 hours | CPU | Protocol 1 | 1:3 | N/A
|  |  | Battery life of up to 12 hours | CPU | Protocol 2 | 0:3 | small sample, no failures allowed
|  |  | Safe for up to 12 hours of wear | CPU | Protocol 1 | 0:3 | small sample, no failures allowed
|  |  | Memory for ≥24 hours of data | CPU | Protocol 2 | 0:3 | small sample, no failures allowed
|  |  | Device will turn on | CPU | Protocol 3 | 1:5 | N/A
|  |  | Device will indicate recording | CPU | Protocol 3 | 1:10 | N/A
|  |  | Device willl charge | CPU | Protocol 3 | 1:5 | N/A
5.12.3 | Safety Requirements | 5.1.2 | Device will not overheat | CPU | Protocol 3 | 1:3 | N/A
|  |  | Apnoeic Event False Positive ≤ 12.5% | Pulse Ox Analysis | Protocol 14 | Pass/Fail (N=25)
|  |  | Apnoeic Event False Negative ≤ 5% | Pulse Ox Analysis | Protocol 14 | Pass/Fail (N=60 )
|  |  | Soft Palate Snoring False Positive ≤ 12.5% | Snoring Analysis | Protocol 15 | Pass/Fail (N=25)
|  |  | Soft Palate Snoring False Negative ≤ 5% | Snoring Analysis | Protocol 15 | Pass/Fail (N=60 )
|  |  | Back of Tongue Snoring False Positive ≤ 12.5% | Snoring Analysis | Protocol 16 | Pass/Fail (N=25)
|  |  | Back of Tongue Snoring False Negative ≤ 5% | Snoring Analysis | Protocol 16 | Pass/Fail (N=60 )
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol | Acceptance Criteria | Comments
5.12.9 | Physical | 5.1.3 | Meets drawn dimensions | Microphone | Protocol 4 | Pass/Fail (N=1) | N/A
|  |  | Adjustable Headband | Microphone | Protocol 8 | Pass/Fail (N=1) | N/A
|  |  | Meets drawn dimensions | Pulse Ox | Protocol 4 | Pass/Fail (N=1) | N/A
|  |  | Snug fit on finger | Pulse Ox | Protocol 8 | Pass/Fail (N=1) | N/A
|  |  | Meets drawn dimensions | CPU | Protocol 4 | Pass/Fail (N=1) | N/A
|  |  | Adjustable Chest Strap | CPU | Protocol 8 | Pass/Fail (N=1) | N/A
5.13.2 | Patient Interface | 5.1.3 | Ports for inputs, charger and SD card | CPU | Protocol 9 | Pass/Fail (N=1) | N/A
|  |  | LED placement records SpO2 | Pulse Ox | Protocol 3 | 1:3 | N/A
|  |  | Microphone placement records audio data | Microphone | Protocol 3 | 1:3 | N/A
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol | Acceptance Criteria | Comments
5.12.2 | Environments of Use | 5.1.9 | Function at 15-38 C | All Hardware | Protocol 10 | 1:3 | N/A
|  |  | Function at 60-100% humidity | All Hardware | Protocol 11 | 1:3 | N/A
|  |  | Survive 1.5 ±.1 m fall | All Hardware | Protocol 12 | 1:3 | N/A
|  |  | Can withstand 1.5 psi compress | All Hardware | Protocol 12 | 1:3 | N/A
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol | Acceptance Criteria | Comments
5.11.1 | Patient Expectations | 5.1.12 | Data Storage Under Card Capacity | Arduino Code | Protocol 5 | 1:8 | N/A
|  |  | Data Storage Over Card Capacity | Arduino Code | Protocol 5 | Pass/Fail | N/A
|  |  | Data Storage with File Already Present | Arduino Code | Protocol 5 | Pass/Fail | N/A
|  |  | Synchronization of Audio and Pulse Oximeter data | Arduino Code | Protocol 6 | 2:32 | N/A
|  |  | Pulse Oximeter is Sampled at 100 Hz | Arduino Code | Protocol 7 | 1:8 | N/A
|  |  | Audio Data is Sampled at 20 kHz | Arduino Code | Protocol 7 | 1:8 | N/A
|  |  | Audio Data is not clipped between 15 - 70 dB | Arduino Code | Protocol 17 | 1:16 | N/A
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol | Acceptance Criteria | Comments
5.13.1 | Human Factors (Ergonomics) | 5.1.2 | Analysis time less than 1 hour | Clinician GUI | Protocol 18 | 1 failure:50 runs | N/A
| Considerations |  | GUI visualization effective for people | Clinician GUI | Protocol 19 | 1 failure:9 people | N/A
| (Ergonomics) |  | Program run failure | Clinician GUI | Protocol 20 | 1 failure:500 runs | N/A
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol
5.11.1 | Patient Expectations | 5.1.2 | Can be worn up to 12 hours | CPU | Protocol 1
|  |  | Battery life of up to 12 hours | CPU | Protocol 2
|  |  | Safe for up to 12 hours of wear | CPU | Protocol 1
|  |  | Memory for ≥24 hours of data | CPU | Protocol 2
|  |  | Device will turn on | CPU | Protocol 3
|  |  | Device will indicate recording | CPU | Protocol 3
|  |  | Device willl charge | CPU | Protocol 3
|  |  | Data Storage Under Card Capacity | Arduino Code | Protocol 5
|  |  | Data Storage Over Card Capacity | Arduino Code | Protocol 5
|  |  | Data Storage with File Already Present | Arduino Code | Protocol 5
|  |  | Synchronization of Audio and Pulse Oximeter data | Arduino Code | Protocol 6
|  |  | Pulse Oximeter is Sampled at 100 Hz | Arduino Code | Protocol 7
|  |  | Audio Data is Sampled at 20 kHz | Arduino Code | Protocol 7
|  |  | Audio Data is not clipped between 15 - 70 dB | Arduino Code | Protocol 17
5.12.3 | Safety Requirements | 5.1.2 | Device will not overheat | CPU | Protocol 3
|  |  | Apnoeic Event False Positive ≤ 12.5% | Pulse Ox Analysis | Protocol 14
|  |  | Apnoeic Event False Negative ≤ 5% | Pulse Ox Analysis | Protocol 14
|  |  | Soft Palate Snoring False Positive ≤ 12.5% | Snoring Analysis | Protocol 15
|  |  | Soft Palate Snoring False Negative ≤ 5% | Snoring Analysis | Protocol 15
|  |  | Back of Tongue Snoring False Positive ≤ 12.5% | Snoring Analysis | Protocol 16
|  |  | Back of Tongue Snoring False Negative ≤ 5% | Snoring Analysis | Protocol 16
5.12.9 | Physical | 5.1.3 | Meets drawn dimensions | Microphone | Protocol 4
|  |  | Adjustable Headband | Microphone | Protocol 8
|  |  | Meets drawn dimensions | Pulse Ox | Protocol 4
|  |  | Snug fit on finger | Pulse Ox | Protocol 8
|  |  | Meets drawn dimensions | CPU | Protocol 4
|  |  | Adjustable Chest Strap | CPU | Protocol 8
5.13.2 | Patient Interface | 5.1.3 | Ports for inputs, charger and SD card | CPU | Protocol 9
|  |  | LED placement records SpO2 | Pulse Ox | Protocol 3
|  |  | Microphone placement records audio data | Microphone | Protocol 3
Design | Input | Design | Output | Verification | Testing
MR Section | Functional Req. | DS Section | Design Req Description | Element | Test Protocol
5.12.9 | Physical | 5.1.3 | Meets drawn dimensions | Microphone | Protocol 4
|  |  | Adjustable Headband | Microphone | Protocol 8
|  |  | Meets drawn dimensions | Pulse Ox | Protocol 4
|  |  | Snug fit on finger | Pulse Ox | Protocol 8
|  |  | Meets drawn dimensions | CPU | Protocol 4
|  |  | Adjustable Chest Strap | CPU | Protocol 8
5.13.2 | Patient Interface | 5.1.3 | Ports for inputs, charger and SD card | CPU | Protocol 9
|  |  | LED placement records SpO2 | Pulse Ox | Protocol 3
|  |  | Microphone placement records audio data | Microphone | Protocol 3
5.12.2 | Environments of Use | 5.1.9 | Function at 15-38 C | All Hardware | Protocol 10
|  |  | Function at 60-100% humidity | All Hardware | Protocol 11
|  |  | Survive 1.5 ±.1 m fall | All Hardware | Protocol 12
|  |  | Can withstand 1.5 psi compress | All Hardware | Protocol 12
5.13.1 | Human Factors (Ergonomics) | 5.1.2 | Analysis time less than 1 hour | Clinician GUI | Protocol 18
| Considerations |  | GUI visualization effective for people | Clinician GUI | Protocol 19
| (Ergonomics) |  | Program run failure | Clinician GUI | Protocol 20

## Sheet2
Element | Test and Equipment | Specifications | Acceptance Criteria
Arduino Data Capture | Data Storage Under Card Capacity | Arduino will record audio and pulse oximeter absorbance data. | Audio and pulse oximeter events are accurately captured.
| Data Storage Over Card Capacity | If the recorded data fills the capacity of the SD card, the recording should be terminated. | Data recording is haulted and no data is overwritten.
| Data Storage with File Already Present | If data is already present on the SD card, the new data will be stored on a separate file. | A new file is created if files already exist on the SD card.
| Synchronization of Audio and Pulse Oximeter data | Audio and pulse oximeter events should occur within 5 seconds of each other to ensure accuracy of analysis. | Recorded audio and pulse oximeter events are observed within 5 seconds of when they actually occurred.
| Pulse Oximeter is Sampled at 100 Hz | Pulse oximeter data will need to be recorded with a minimum frequency of 80 Hz to ensure proper analysis. | > 100 Hz Sampling Frequency
| Audio Data is Sampled at 10 kHz | Snoring Data will need to be recorded with a minimum frequency of 10 kHz to ensure proper analysis. | > 10 kHz Sampling Frequency
| Audio Data is not clipped between 15 - 70 dB | Recorded audio values between 15-70 dB (the dynamic range) will not be clipped. | Data in the dynamic range doesn't exhibit distortion from being clipped.

## Sheet3
Data Analysis | What is this sheet for?
Element | Spec | Acceptance Criteria | Test Method/ Protocol
Apnea Data Analysis Code | False positive detection of apnoeic event 12.5% (5.1.2)